ACME performance and precision with multiple Powercapes
using HWMON/Sigrok

---
Measurement principle
Measure the time from update to update for the first probe @0x40 and verify there is always enough time to read out all registers until next sample time is due
Write trace “a” when entering update_device
Write trace “b” after completion of i2c operations
Measure time bk– ak	cost of i2c transfer
Measure time bk – bk-1	achievable sampling period
Do this for both versions of the hwmon driver

---
Test commands
Data acquisition
trace-cmd start -p nop
sigrok-cli -d baylibre-acme --samples 200 --config samplerate=${freq}
trace-cmd stop
trace-cmd extract
Possible frequency values with sigrok-cli
Libsigrok/src/hardware/baylibre-acme will set the update interval through sysfs, possible values are discrete however, resulting of the possible averaging parameter times the constant Integration Time setting of 2.2ms
Averaging values	1	4	16	64	128	512	1024
Periods (s)	0.0022	0.0088	0.0352	0.1408	0.2816	1.1264	2.2528
Freq (Hz)	454	113.63	28.41	7.1	3.55	0.887	0.443
Tested (Hz)	454	113	28	7
Visualization
Done with kernelshark

---
Preliminary discussion
Expected achievable sample rates
With the 4.3 hwmon driver, upon accessing sysfs with the show function, all 8 registers are read once per programmed update interval. The 4.4 driver will read only the required value register.
Reading all registers takes at best 500us * 8 = 4ms per probe, but this can be much longer if the threaded IQR completion is delayed (i2c-omap).
The table bellow provides the best sampling frequency (Hz) based on the number of connected probes, and the driver version, where no old value is repeat.
#probes / #channels	1 Probe	2 Probes	3 Probes	4 Probes	5 Probes	6 Probes	7 Probes	8 Probes
Driver 4.3 / Any ch #	113	113	28	28	28	28	28	28
Driver 4.4 / 1 channel	454	454	454	454	113	113	28	28
Driver 4.4 / 2 channels	454	454	454	113	113	113	28	28
Driver 4.4 / 3 channels	454	454	113	113	113	28	28	28
Driver 4.4 / 4 channels	454	454	113	113	28	28	28	28

---
4.3 Driver, 8 probes, 7 Hz
Command and driver setup trace verification
sigrok-cli -d baylibre-acme --samples 200 --config samplerate=7
[ 4232.068759] Setup averaging = 64	64*2.2ms = 141ms ~ 7.1 Hz OK
[ 4232.072163] Actual interval = 140800 us
Results
Elapsed time between sigrok-cli readout:	~200 ms	KO, see [1]
Elapsed time between fresh values:	~200 ms	KO, see [2]
Comment
[1] with driver 4.3 bl_acme_receive_data requires a small patch to prevent introducing an additional delay.
[2] driver 4.3 will prevent the read out every second time, this might be related to how timer events for sleeping are managed.

---
The remaining tests and measurements are done with hwmon driver v4.4 (the one using regmap) because v4.3 is even worse, and linux-next will provide v4.4 anyhow.

---
4.4 Driver, 8 probes, 28 Hz
Command and driver setup trace verification
sigrok-cli -d baylibre-acme --samples 200 --config samplerate=28
[ 39.494589] Setup interval = 35 us
[ 39.498304] Setup averaging = 16
[ 39.501685] Actual interval = 35200 us	OK
Results
Elapsed time between fresh values:	~37 ms	(sigrok timer)	OK
Comment
Kernelshark shows very regular readout times for probe0/vshunt, meaning that this configuration can be sustained in terms of performances.

---
4.4 Driver, 8 probes, 113 Hz
Command and driver setup trace verification
sigrok-cli -d baylibre-acme --samples 200 --config samplerate=113
[ 4418.708234] Setup interval = 8 us
[ 4418.711755] Setup averaging = 4
[ 4418.715047] Actual interval = 8800 us	OK
Results
Elapsed time between fresh values:	32 ms	(sigrok timer)	KO
Comment
As expected, setting a readout period under the required time spent in i2c-omap will result in an erratic sample rate (bottleneck effect).
Sample rate above 28 Hz result in time stretching (repeated samples).

---
Conclusions
Using 8 probes is functional, but with sample freq restrictions (28 Hz max) for accurate signal properties.
Recommended changes in Current End-user package
It is important to provide users with:
the version of sigrok that uses the timer device.
the hwmon driver 4.4, as it will allow for better sampling frequencies for simple use-cases (see table in slide 4)
Changes in Documentation
I recommend providing table 4 (remove the driver-4.3 line) to the user, with explanation
Using higher sampling rates is always possible to avoid a slow GUI refresh rate, but any sampling freq above the compatible one will result in time stretching (repeated samples).
Signal process software can probably do a better job with proper interpolation algos.
Corrective action and evolutions
Support smaller integration times with hwmon than the fixed 2.2ms to increase sample freq, and switch to IIO driver when sigrok pluging is available (no sysfs)
Investigate using a bitclock higher than 400kHz, might not be possible with the current gpio routers ? Might be difficult to upstream
Investigate using a non-threaded irq instead of a threaded_irq + completion routine for i2c-omap. Might be difficult to upstream.